## Slide 1: The evaluation of JMA’s standard gas scale for measurements of atmospheric methane
*Shuichi Hosokawa1, Teruo Kawasaki1, Yuji Esaki1, Atsushi Takizawa1, Shinya Takatsuji1, Susumu Hashimoto1, Kohshiro Dehara1,
Kentaro Kozumi1,Hidekazu Matsueda2, Yousuke Sawa2,
Kazuhiro Tsuboi2 and Yosuke Niwa2
1. Japan Meteorological Agency, Tokyo, Japan,
2. Meteorological Research Institute, Ibaraki, Japan

## Slide 2: Today’s Topic
（１） CH4 calibration method and standards in JMA
・JMA calibrates CH4 with GC/FID (Gas Chromatography with Flame Ionization Detector) using by 5 standard gases.
・JMA has used 2 Generation series (1G and 2G) of standard gases since 2006.
（２）Long-term Activity of JMA/WCC
・ JMA manage World Calibration Center for CH4 in Asia and the South-West Pacific (WCC), and has carried out reference gases inter-comparisons (Round Robin) object to identify the differences of their standard scales as well as to monitor the long-term stability of standard gases.
（３）Pre-study for replacing JMA calibration system with CRDS
・JMA plans to replace the CH4 calibration system to CRDS
(Cavity Ring Down Spectroscopy) .
・In advance, we tested of GC/FID and CRDS.

## Slide 3: (1) JMA CH4 Calibration System with GC/FID
Matsueda et al., PMG [2004]
※GAW Compatibility Goal : ±2ppb

## Slide 4: JMA’s CH4 Standard gases

## Slide 5: Traceability of JMA Standard 1G→2G
D = Difference ( Measured – Assigned )[ppb]
D =0.96
D =0.41
D =0.27
D =0.73
D =0.47
: NOAA calibration
: Average of concentration
measured by 1G

## Slide 6: Long-term stability of JMA/1G standards
Tsuboi et al., PMG [2016]

## Slide 7
(2) JMA/WCC Round Robin and iceGGO
You can see the result of JMA/WCC Round Robin.
http://ds.data.jma.go.jp/gmd/wcc/ch4/com_annex2.html
JMA/WCC
Round Robin
IceGGO
Inter-Comparison Experiments for Greenhouse Gases Observation

## Slide 8: JMA/WCC Round Robin and iceGGO
※GAW Compatibility Goal : ±2ppb

## Slide 9: JMA/WCC Round Robin and iceGGO
Kawasaki et al., GGMT[2015]
※GAW Compatibility Goal : ±2ppb

## Slide 10: Converted each scale to NOAA
● CSIRO94 and NIST94 scale converted to
NOAA04 scale using by Dlugokencky E.J.
et al. , 2005, and WMO DATA SUMMARY.
● TU/MRI/AIST/NIES/NIPR each scale and
NOAA2004 scale have linear relationship.
→Convert every data to NOAA04 scale.
Tsuboi et al., PMG [2016]
Example for iceGGO experiments

## Slide 11: JMA/WCC Round Robin and iceGGO
※GAW Compatibility Goal : ±2ppb

## Slide 12: （３）Planning to replace of calibration system
JMA plans to replace the calibration system
from GC/FID to CRDS.
Schematic diagram of
the methane calibration system (plan)

## Slide 13: GC/FID vs CRDS 〜 Laboratory Experiment 〜
Measure some cylinders by GC/FID and CRDS
Relationship of these methods is linear.
88% of these difference is within ±2ppb.
※GAW Compatibility Goal : ±2ppb

## Slide 14: Conclusion
The stability and traceability of JMA’s CH4 standard gases have been ensured.
Converting scales, JMA/WCC activity and iceGGO is consistent with each other.
We tested the difference of GC/FID and CRDS. We will operate much comparing tests with actual analyzer and replace the calibration system from GC/FID to CRDS carefully.

## Slide 15: Reference
Tsuboi et al.(2016), Scale and stability of methane standard gas in JMA and comparison with MRI standard gas; Papers in Meteorology and Geophysics Vol.66 15-24, doi: 10.2467/mripapers.66.15
Kawasaki et al.(2016), Inter-comparison experiments of standard gases for JMA/WCC activity; GAW Report 104-109
Matsueda et al.(2004), Methane standard gases for atmospheric measurements at the MRI and JMA; Papers in Meteorology and Geophysics Vol.54 91-113
Dlugokencky et al. (2005): Conversion of NOAA atmospheric dry air CH4 mole fractions to a gravimetrically prepared standard scale, J. Geophys. Res., 110, D18306, doi:10.1029/2005JD006035.